# LABORATORY DECOMMISSION CHECKLIST 2019

Principal Investigators (PI) or Laboratory Managers (LM) relocating, vacating or closing laboratory areas are responsible for leaving laboratories in a suitable condition for re-occupancy or renovation.  If the decommission process is not completed by the PI or LM, responsibility for completion of this process will fall to the associated department.  The Finance Department will continue to charge the responsible department for rent until all applicable checklist items have been completed and reviewed by the Occupational Safety Department. The Radiation Safety Officer and the Institutional Biosafety Officer will participate in the review if applicable.

Complete the following laboratory close-out checklist by checking YES if the applicable item has been completed or N/A if it does not apply.  After completing the checklist the PI or LM must then sign this document indicating that the area(s) listed below are suitable for re-occupancy.  Copies of the completed close-out document must be provided to the applicable departments indicated at the end of this document.  The Occupational Safety Department (OSD) must then be contacted for a final exit inspection of your area(s).

If there are plans to transfer equipment, supplies, reagents, specimens, etc., from the current occupant to the new occupant, and the new occupant has agreed to accept those materials, this must be documented in a separate memo and attached to this document.  A brief description of the items being transferred and a signature from both parties involved indicating they are in agreement is required.

List the laboratory area(s) that will be included in the close-out process.  Include Department name, building name and room numbers.

________________________________________________________________________

________________________________________________________________________

________________________________________________________________________

# GENERAL:

# YES 	 N/A

Decommissioning of the laboratory must only be conducted by authorized laboratory personnel that are trained on the hazards of the lab and the materials contained in it (i.e., Hazard Communication Training, Bloodborne Pathogens Training, Laboratory Safety Training, Laboratory Specific Training, etc.).

# YES 	 N/A

Hazardous materials from laboratory equipment must be identified for disposal before laboratory equipment is discarded.  OSD can assist with the identification process. This may include batteries, transformers, mercury switches, mercury thermometers, mercury lamps, oils, asbestos linings, asbestos jackets, asbestos gloves, radioactive sources, lead shielding and CFC’s from refrigerators/freezers, etc.

Note#1: MCE can remove refrigerants from some refrigerators/freezers. Create a Service Request in the online 2-FIXX portal: https://eam.inforcloudsuite.com/web/base/portallogindisp?tenant=RUSH_PRD

Midwest Lab Services (847) 215-0200 must be contracted to remove refrigerants from ultra low freezers.

Note#2: Refrigerators and freezers must not be left behind in any lab areas unless agreed to in writing between the parties involved. There are expenses associated with refrigerant removal and disposal.  This expense is not the responsibility of future occupants or the institution.

# YES 	 N/A

Before disposing, moving or transferring any equipment reference the Corporate Finance > Fixed Assets website on the Rush intranet. When equipment is transferred to another department, moved or disposed of, the online Asset Disposal / Move / Transfer Form should be completed.

# YES 	 N/A

Internal moves (within Rush) or disposal of items and equipment can be arranged online through the 2-FIXX portal by selecting the MCE Movers category.

# YES 	 N/A

Laboratory equipment such as incubators, shakers, centrifuges, microscopes, sequencers, scintillation counters, etc. must not be left behind.

# YES 	 N/A

Glassware and other glass must be disposed of in glass disposal boxes intended for this purpose or packaged in sturdy cardboard with newspaper or other packing material to prevent breakage.  The boxes must then be labeled glass for disposal.

# YES 	 N/A

Garbage must not be left behind.  All cabinets and drawers must be emptied.

# YES 	 N/A

Common areas such as dark rooms, storage rooms, cold rooms, flammable cabinets and shared labs must be accounted for when addressing any of the items on this checklist.

# CHEMICALS AND HAZARDOUS WASTE:

# YES 	 N/A

Controlled substances must be disposed of or transferred as specified by the Drug Enforcement Agency (DEA) permit under which they are held.

# YES 	 N/A

All laboratory chemicals, chemical containers, chemical wastes and cleaning chemicals must be removed from the laboratory.

# YES 	 N/A

All hazardous waste, unwanted chemicals and empty chemical containers must be disposed of through the Occupational Safety Department (OSD).

Note#1: OSD reserves the right to charge back disposal costs for chemicals that are extremely expensive to dispose of and/or major laboratory clean-outs.

Note#2: Laboratory personnel are responsible for filling out a Requisition for Chemical Disposal (attachment to the Chemical Waste Management Policy #0302) on the Rush intranet and arranging for a hazardous waste pick-up.

# YES 	 N/A

Storage cabinet areas must be clean; chemical residues, spills and drips, must be cleaned and decontaminated.

# YES 	 N/A

Bench top liners and/or disposable covers must be removed from work surfaces.

# YES 	 N/A

Laboratory bench tops must be washed with soap and water.

# YES 	 N/A

All materials must be removed from chemical fume hoods.  All surfaces in the fume hood that can be reached must be cleaned and decontaminated.

# YES 	 N/A

All materials must be removed from sinks and sinks must be cleaned.

# YES 	 N/A

All unknown wastes must be identified with proper labeling (chemical name and hazards).

# YES 	 N/A

All chemicals must be removed from shared storage areas.  This includes common areas such as refrigerators, freezers, cold rooms, flammable cabinets, corrosives cabinets, other people’s labs, chemical fume hoods and dark rooms.

# YES 	 N/A

If hazardous chemicals are to be shipped to another facility (off Rush campus) a permitted hazardous materials handler/hauler must be contracted to legally transport the hazardous materials.  OSD should be contacted for assistance with obtaining a hazmat handler/hauler.  If you ship the chemicals yourself, you must have a current Dangerous Goods training certification and ship the materials according to DOT/IATA regulations.

# YES 	 N/A

All chemical hazard warning labels must be removed from bench tops, doors, cold rooms, storage areas, etc.  The lab hazard identification sign posted outside the lab must be removed also.

# YES 	 N/A

Medical Center Engineering must be contacted to shut off medical gas line supplied oxygen and natural gas.

# RADIOACTIVE MATERIALS:

# YES 	 N/A

The Radiation Safety Department must be contacted to decommission the laboratory.

# YES 	 N/A

All radioactive materials/sources must be removed from the laboratory for proper disposal through the Radiation Safety Department.

# YES 	 N/A

All radiation warning labels must be removed from bench tops, hoods, doors, storage areas, lead shielding, etc.

# Note: Labels should only be removed at the discretion of Radiation Safety.

BIOHAZARDOUS MATERIALS AND SHARPS-                                        REGULATED MEDICAL WASTE (RMW):

# YES 	 N/A

All biohazardous waste materials must be packaged in red bag lined Stericycle RMW containers.  This includes human and animal tissue, blood and body fluids, cultures, cell lines, DNA, bacteria, viruses, prions, etc.

Note#1: When disposing of pathological wastes from humans or animals the Stericycle RMW container must be labeled with a For Incineration Only sticker which can be obtained from Environmental Services.

Note#2: All red bags containing biological liquids must have enough absorbent material to prevent a spill if the bag breaks inside the Stericycle RMW container.

Note#3: If you need empty Stericycle RMW containers dropped off or full containers removed, contact Environmental Services at 25290. Laboratory employees are responsible for arranging for the pickup of full containers.

# YES 	 N/A

Slides containing tissue samples must be disposed of in sharps containers.

# YES 	 N/A

All used and unused sharps (blades, scalpels, glass pipettes, glass capillary tubes, needles, syringes) must be disposed of in sharps containers. If you need additional sharps containers or full containers picked-up, page the Biosystems Technician at 85-7238 or 85-5634.

Note: Biosystems only pick up recyclable sharps containers. These can be identified by the Stericycle name and logo on the container as seen below.

If the sharps container is a disposable one time use container, snap the container 	shut and it can be disposed of in the red bag medical waste container.

# YES 	 N/A

If tissue samples are stored in a liquid preservative, the liquid is typically a hazardous chemical waste and must not be disposed of in the sanitary sewer or in the Stericycle RMW container.  The tissue sample must be separated from the preservative before disposal in the Stericycle RMW container.  The resulting liquid preservative must be collected for hazardous chemical waste disposal.

# YES 	 N/A

Defrost, clean and decontaminate refrigerators, freezers and any other equipment that was used to handle or store biological material.  Use a 10% bleach solution or hospital approved disinfectant such as Virex TB for decontamination.

# YES 	 N/A

All laboratory surfaces that have come in contact with biological materials must be cleaned and decontaminated.  Use a 10% bleach solution or hospital approved disinfectant such as Virex TB for decontamination.

# YES 	 N/A

All biological safety cabinets must be completely decontaminated by the RUMC approved vendor if you are leaving the cabinets in the lab or if you are taking them with you. If you move the cabinet to a new location it must be recertified after the move.

Note#1: Future occupants, construction workers, equipment movers, etc. do not know what types of biological materials were used in biological safety cabinets that are left behind.  Therefore, it is imperative that all biological safety cabinets be completely decontaminated by an approved and qualified outside contractor.  For a list of approved contractors contact OSD at 27233.  This expense is not the responsibility of future occupants or the institution.

Note#2: If the PI or LM knows the history of the biological safety cabinet does not involve the use of any infectious agents or biological materials hazardous to humans, an exception to complete decontamination by an outside contractor can be made.  For this exception to be accepted for lab close-out purposes the PI or LM must attach a letter in writing with this close-out indicating that no such infectious agents or biological materials hazardous to humans have been used in the cabinet in question and that the cabinet is safe to be handled.  If there is more than one biological safety cabinet this must be indicated in the letter by the serial numbers.

Contact the Rush Institutional Biosafety Officer 2-5222 for questions.

# YES 	 N/A

All media and supplies must be removed or disposed of.

# YES 	 N/A

All universal biohazard symbols and biohazard warnings must be removed from refrigerators, freezers, doors and other areas within the lab.

# YES 	 N/A

Unwanted red biohazard bags, orange autoclave bags and biohazard labeled specimen bags and boxes must not be disposed of in the regular trash.  These items must be disposed of in the Stericycle RMW containers.

# YES 	 N/A

If biological materials are to be shipped to another facility (off Rush campus) a permitted hazardous waste handler/hauler must be contracted to legally transport the hazardous materials or you must ship these materials in compliance with the Rush Hazardous Materials/Dangerous Goods Shipment Policy Number OP-0296.  OSD should be contacted for assistance with obtaining a waste handler/hauler.

# COMPRESSED GAS CYLINDERS:

# YES 	 N/A

Respiratory Care (26384) or the appropriate vendor must be contacted to remove all compressed gas cylinders.

Note: Laboratory personnel are responsible for arranging lab access for Respiratory Care or the vendor and ensuring the cylinders are removed.

# YES 	 N/A

All regulators must be removed and valve protection caps securely placed.

# YES 	 N/A

All excess regulators, valve protection caps and valves should be returned to Respiratory Care.

# YES 	 N/A

The Medical Center compressed gas cylinder vendor should be contacted directly to move or remove liquid nitrogen tanks.

# MOVING FROM ONE LABORATORY TO ANOTHER ON CAMPUS:

# YES 	 N/A

All hazardous materials must be moved by authorized laboratory personnel that are trained in handling such materials (i.e., Hazard Communication Training, Bloodborne Pathogens Training, Laboratory Safety Training, etc.).  Consult with Radiation Safety about moving radioactive materials.

# YES 	 N/A

Hazardous chemicals must be segregated into compatible hazard classes before transporting them to the new lab.  The different hazard classes of chemicals must be separated by secondary spill containment during transport.

# YES 	 N/A

Chemicals must be transported on a cart with a lip or sump.  If you need to borrow a cart contact Occupational Safety at 27233.

# YES 	 N/A

Spill kit materials must accompany any liquid hazardous materials transported through the Medical Center.

# YES 	 N/A

If you have a chemical spill and assistance is needed contact OSD at 27233 or pager 85-3503.

# YES 	 N/A

Proper personal protective equipment must be available during the transport of hazardous materials through the Medical Center.  Do not press elevator buttons, door knobs, handrails, etc. with gloved hands.  If you must, wear a glove on one hand keeping the other hand free to manipulate doors, elevator buttons, etc.

# YES 	 N/A

Respiratory Care should be contacted to move compressed gas cylinders.

# YES 	 N/A

The laboratory hazard identification sign located outside the old lab must be removed and a new sign must be filled out and posted outside the new laboratory.

# YES 	 N/A

All mercury containing thermometers and mercury compounds must be disposed of through OSD before moving to the new laboratory.

Your signature indicates that all applicable checklist items have been completed and the indicated areas are suitable for re-occupancy or renovation.  Contact Occupational Safety 27233 for the final exit inspection.

# Principal Investigator Signature: __________________________________________

Date: _________________

Distribute copies of this completed checklist to:

- Occupational Safety (FAX 24229) or OccSafety <OccSafety@rush.edu>

- Your Department Head/Administrative Officer

- Parisa Ramsi, Research Administrator, for research labs only

- Allison Cantu, RML Administrative Director, for RML only